| Roles | Name | Designation | Signature | Date |
| --- | --- | --- | --- | --- |
| Author |  |  |  |  |
| Reviewer |  |  |  |  |
| Approver |  |  |  |  |

| Revision History |  |  |  |  |
| --- | --- | --- | --- | --- |
| DCO no. | Revision | Change Description | Author | Effective Date |
| DCO XXXXX | 01 | Initial Release |  | XX-XXX-XXXX |

| Project code/ number |  | Date |
| --- | --- | --- |
| Product Description |  |  |
| Intended Use |  |  |
| Instructions for use |  |  |

| Project code/ number | SickleFind |  |  | Date |  |
| --- | --- | --- | --- | --- | --- |
| Product Description |  |  |  |  |  |
| Intended Use |  |  |  |  |  |
| Instructions for use |  |  |  |  |  |
| Functional Requirements |  |  |  |  |  |
| Sl no | Input | Specification | Output |  | Verification |
| 2 |  |  |  |  |  |
| 3 |  |  |  |  |  |
| 4 |  |  |  |  |  |
| 5 |  |  |  |  |  |
| Performance Requirement |  |  |  |  |  |
| Sl no | Input | Specification | Output |  | Verification |
| Usability requirement |  |  |  |  |  |
| Sl no | Input | Specification | Output |  | Verification |
|  | Minimal training requirement, in a clinical laboratory testing center. | Only pipetting skills are required. | Pipetting skills are a prerequisite to follow the instruction manual. |  | Training was given to operator prior to the validation tests |
| Safety requirement |  |  |  |  |  |
| Sl no | Input | Specification | Output |  | Verification |
|  | Non Hazardous components included | Hazardous Components like phenol, chloroform not included in the formulation Instruction manual includes safety instructions | Hazardous chemicals not included. See MSDS of chemicals Instruction manual includes safety instructions |  | Refer to MSDS Refer to IFU |
| Design considerations |  |  |  |  |  |
| Sl no | Input | Specification | Output |  | Verification |
|  | Usability | Minimal reagents, minimal steps and shorter turn-around time | Refer to instruction manual. Test only requires small quantities of reagents. The test has a simple workflow and can be performed in 15 minutes. |  | Refer to IFU. Test can be performed in 15 minutes. Kit contains 3 tubes that can be easily mixed together to prepare the working buffer. The results can be obtained in 15 minutes. |
|  | Compatibility | Reagents used should be compatible with blood samples | Reagents used are compatible with blood samples. Refer to MSDS |  | Reagents are compatible with blood samples. |
|  | Transport ease | All the reagents and kit components used should be stable at ambient temperature which can be shipped without a cold chain. | The kit is stable at room temperature and can be shipped without a cold chain. |  | The kit is stable at room temperature No Special storage condition. |
|  | Safety | Reduce hazardous chemicals | No hazardous chemicals were used in the kit preparation. |  | The kit does not contain hazardous chemicals. Refer to MSDS |
| Regulatory Considerations |  |  |  |  |  |
| Sl no | Input | Specification | Output |  | Verification |
|  | Kits should be manufactured as per the requirements of ISO 13485, ISO 14971 CDSCO Medical Device Regulations (MDR) | ICMR guidelines and MDR guidelines should be followed | ICMR and MDR guidelines were followed. Test license approval received from CDSCO prior to clinical validation and analytical performance tests. |  | Clinical testing studies and analytical performance studies are being conducted at ICMR Jabalpur. |
|  | IFU, Labelling to be prepared as per the requirements of ISO 15223-1:2016 | ISO 15223-1:2016 Medical devices — Symbols to be used with medical device labels, labelling and information to be supplied — Part 1: General requirements | IFU, labelling were prepared as per requirements of ISO 15223-1:2016 |  |  |
| Other inputs |  |  |  |  |  |
| Sl no | Input | Specification | Output |  | Verification |
|  | Shelf life |  |  |  |  |
|  | Low cost |  |  |  |  |
|  | Storage |  |  |  |  |
|  | Handling Instruction | The technicians are expected to follow the safety protocols followed while handling blood samples, i.e wearing gloves and other PPE. | Safety protocols and handling instructions are included in IFU. |  | Refer IFU |