| Roles | Name | Designation | Signature | Date |
| --- | --- | --- | --- | --- |
| Author |  |  |  |  |
| Reviewer |  |  |  |  |
| Approver |  |  |  |  |

| Revision History |  |  |  |  |
| --- | --- | --- | --- | --- |
| DCO no. | Revision | Change Description | Author | Effective Date |
| DCO XXXXX | 01 | Initial Release |  | XX-XXX-XXXX |

- PURPOSE

To establish procedures and guidelines for product design and development. This procedure serves to:

1.1 Control the design of the device in order to ensure that specified design requirements are met.

1.2 Guide development personnel in systematically designing and developing new products.

1.3 Facilitate proper planning.

1.4 Facilitate proper documentation and record retention.

1.5 Define the responsibility of all functions and personnel involved in development.

- SCOPE

Applies all products that are all fully or partially designed and developed by Kyron Healthcare.

- REFERENCES

  - External Documents

    - EN ISO 13485:2016

  - Related Documents

    - Quality Manual

    - QS-03-QA Control of Documents and Record Procedure

    - QS-06-QA Risk Management Procedure

    - QS-07-QA Process Validation and Process Control Procedure

    - QS-11-QA Engineering Change Order Procedure

- DEFINITIONS

| DHF | Design History File, means a compilation of records which describes the design history of a finished device |
| --- | --- |
| Design Input | The physical and performance requirements of a device that are used as a basis for device design. This includes but not limited to regulatory and safety requirements. A design effort is required to translate these design inputs into design outputs. |
| Design Output | The results of a design effort at each design phase, leading to the finished design output. The finished design output is the basis for the device master record. The total finished design output consists of the device, its packaging and labelling, and the device master record |
| Design Review | A documented, comprehensive, systematic examination of a design to evaluate the adequacy of the design requirements, to evaluate the capability of the design to meet these requirements, and to identify problems |
| DMR | Device Master Record, a compilation of records containing the procedures and specifications for a finished device. |

- RESPONSIBILITY

  - Product Development Department is responsible for the contents of this procedure.

  - The project team will carry out all design and development activities of the new product. Depending on the nature of the project, the team may consist of one of more representatives from the following groups that can provide valuable and necessary information for market analysis, technical analysis, quality and regulatory compliance, design inputs as well as support on assembly, experiment and inspection on prototypes during all stages of design and development:

- Product Development Department

- Sales and Marketing Department

- Quality Assurance / Regulatory Affairs Department

- Manufacturing/ Production Department

- Other Departments

  - Product Development Department is responsible for the following:

- Maintain project documentation, which includes all drawings, reports and records.

- Initiates purchase requisition for materials on prototypes and tooling.

- Provides vendors or production with the necessary design, drawings and specifications.

- Identifies all input requirements relating to the product, including applicable statutory and regulatory requirements.

- Schedules design review meeting to verify completion of assigned tasks and file design review documents.

- Ensures that the design outputs are verified to meet design input requirements during all stages of design and development.

- Ensures that the design outputs are validated to meet user needs.

- Ensures that design transfer activities take place successfully.

- PROCEDURES

  - Project Initiation

    - Management appoints a Project Manager & members to form a project team to complete each product development and design control project. The project team has to fill the Project Initiation Form (FM-05001). The Project Initiation Form contains the purpose of starting the project and records the responsibilities and authorities of each team member.

    - The resources needed as well as the competence of the personnel shall be indicated in this record

  - Design and Development Planning

    - There shall be a plan to describe or reference the design and development activities and define responsibility for implementation. Time and resources required for each of the activities shall be included.

    - Product Development Schedule or GANTT chart shall be used to plan and track the design and development stages which consist of but not limited to the following:

      - Design Input

      - Design Output

      - Design Review

      - Design Verification

      - Design Validation

      - Design Transfer

      - Design Change

      - Regulatory Planning

      - Post-Market Monitoring

    - The plan shall be reviewed during Design Review. It shall be updated to reflect the actual status of the activities as the design and development evolves.

  - Design and Development Inputs

    - Design inputs would include:

- Functional, performance, usability and safety requirements, according to the intended use;

- Applicable regulatory requirements and standards;

- Applicable output(s) of risk management;

- As appropriate, information derived from previous similar designs;

- Other requirements essential for design and development of the product and processes.

    - These inputs shall be recorded in the Design Input Form (FM-05001) and reviewed for adequacy and approved by the project team.

    - These inputs shall be complete, unambiguous and able to be verified or validated and not in conflict with each other.

    - These inputs shall be quantitative whenever possible and include ranges where appropriate.

  - Design & Development Outputs

    - Products are designed to meet the requirement of the design inputs.

    - Design output shall provide, where applicable, appropriate information for purchasing, production, and service provision, which may be in the form of drawings, specifications (including product acceptance criteria) It shall also contain characteristics in terms of product design or labelling that are essential for its safe and proper use.

    - Design output requirements shall be verified and recorded in the Design Output Form (FM-05002)

    - The project team shall review and approve all design output requirements before releasing this for prototype or mass production.

  - Design and Development Review

    - The Design Review shall take place in accordance with the Product Development Schedule (FM-05002)

    - Design Review is conducted to evaluate the ability and feasibility of design outputs to meet design inputs and to identify and propose necessary actions.

    - Records of Design Review shall be maintained in Design Review Minutes (FM-05006).

    - Depending on the nature of the review, not all members of the project team must be present. However, the specific personnel must be present if the personnel’s area of responsibility would be reviewed.

    - If certain personnel are absent from the design review, information on the review would be passed to these personnel in the form of an email. These same personnel would then need to sign on the Design Review Minutes to indicate that he or she has acknowledged the contents of the review.

  - Design and Development Verification

    - Design verification shall be performed in accordance with the Product Development Schedule to ensure the design outputs meet the design inputs via Design Input & Output Matrix Form (FM-05005)

    - Test Protocol (FM-07001) shall be drafted and approved before actual design verification can take place. These plans shall minimally contain information such as methods, acceptance criteria and statistical techniques with rationale for sample size.

    - Test Report (FM-07002) shall indicate the results, analysis of results and conclusion.

    - If the intended use requires that the medical device be connected to, or have an interface with, other medical device(s), verification shall include confirmation that the design outputs meet design inputs when so connected or interfaced.

    - Test Protocol and Test Report containing the results and conclusion of the verification and necessary actions shall be maintained as per Control of Documents and Records Procedure (QS-03-QA).

  - Design and Development Validation

    - Design validation shall be performed in accordance with the Product Development Schedule to ensure that the resulting product is capable of meeting the requirements for the specified application or intended use.

    - Design validation shall be conducted per approved validation plans. These plans shall minimally contain information such as methods, acceptance criteria, and statistical techniques with rationale for sample size.

    - Design validation shall be conducted on representative product. Representative product includes initial production units, batches or their equivalents. The rationale for the choice of product used for the validation shall be recorded.

    - Clinical Evaluation or performance evaluations of the medical device in accordance with applicable regulatory requirements shall be performed. A medical device used for clinical evaluation or performance evaluations is not considered to be released for use to the customer.

    - If the intended use requires that the medical device be connected to, or have an interface with, other medical device(s), validation shall include confirmation that the requirements for the specified application or intended use have been met when so connected or interfaced.

    - Validation shall be completed prior to release for use of the product to the customer unless permitted under special regulatory routes. Refer to Process Validation and Process Control Procedure (QS-07-QA).

    - Validation test protocol (FM-07001) and validation report (FM-07002) containing the results and conclusion of the validation and necessary actions shall be maintained as per Control of Documents and Records Procedure

(QS-03-QA).

  - Design and Development Transfer

    - The finished design outputs form the basis of Device Master Record (DMR) and shall be reviewed and approved accordingly as per Engineering Change Order procedure.

    - The DMR shall be accurately translated into production specifications during design transfer. It consists of the manufacturing procedures, material specifications, drawings, inspection procedures, etc. that are essential to build the finished device.

    - The DMR is transferred to the Production / Contract Manufacturer so that the products can be built in the manufacturing environment i.e. under qualified and validated equipment and facilities, and under validated process and trained personnel with approved procedures.

  - Control of Design and Development Changes

Design changes after product release shall be controlled according to Engineering Change Order procedure. Kyron shall determine the significant of the change to the function, performance, usability, safety and applicable regulatory requirements for the device and its intended use.

Changes shall be controlled by this procedure (QS-05-EG) and is to be documented in the Design Change History form (FM-05007).

  - Design History File

Product Development Department shall establish and maintain a DHF for each type of medical device or medical device family. The DHF shall contain or reference the records generated to demonstrate conformity to the requirements for design and development and records for design and development changes.

- RECORD

7.1 FM-05001   Design Development Plan

7.2 FM-05002   Resource Requirement

7.3 FM-05003   Design Input Form

7.4 FM-05004   Design Output Form

7.5 FM-05005   Design Review Minutes

7.6 FM-05006   Design Change History Form

7.7 FM-05007   Design Verification

7.8 FM-05008   Design Validation

7.9 FM-05009   Design Master File

7.10 FM-05010   Design Transfer

7.11 FM-05011   Design History File

- APPENDIX

NA