| Roles | Name | Designation | Signature | Date |
| --- | --- | --- | --- | --- |
| Author |  |  |  |  |
| Reviewer |  |  |  |  |
| Approver |  |  |  |  |

| Revision History |  |  |  |  |
| --- | --- | --- | --- | --- |
| DCO no. | Revision | Change Description | Author | Effective Date |
| DCO XXXXX | 01 | Initial Release |  |  |

- PURPOSE

  - To define the process of identifying processes that require validation, planning and executing the validation programs, and maintaining records of validation.

  - To ensure that process validation requirements are correctly applied and all aspects required for performing process validation are properly identified and reviewed.

- SCOPE

The scope applies to the manufacturing processes for medical devices and specifies specific recommendations for verification of design output and design validation. This applies to any process in the Company that cannot be fully verified by subsequent inspection and/or test.

- REFERENCES

  - External Documents

3..1.1     ISO 13485:2016

  - Related Documents

    - Quality Manual

    - QS-06-RD  Design Control Procedure

- DEFINITIONS

| Installation qualification (IQ) | Establishing by objective evidence that all key aspects of the process equipment and ancillary system installation adhere to the manufacturer’s approved specification and that the recommendations of the supplier of the equipment are suitably considered. |
| --- | --- |
| Operational qualification (OQ) | Establishing by objective evidence process control limits and action levels which result in product that meets all predetermined requirements |
| Performance qualification (PQ) | Establishing by objective evidence that the process, under anticipated conditions, consistently produces a product which meets all predetermined requirements |
| Process validation | Establishing by objective evidence that a process consistently produces a result or product meeting its predetermined requirements |
| Process validation protocol | A document stating how validation will be conducted, including test parameters, product characteristics, manufacturing equipment, and decision points on what constitutes acceptable test results |
| Verification | Confirmation by examination and provision of objective evidence that the specified requirements have been fulfilled |
| Attribute data | Data collected from an inspection process that results in a go/no-go, or accept/reject decision |
| Variable Data | Data collected from measurements taken on the process output. Variable data can be measured and expressed as a number on a continuous scale of measurement |

- RESPONSIBILITY

The Quality Assurance is responsible for coordinating activities with Manufacturing and R&D to plan and implement activities under this program. All personnel involved in the validation processes shall be trained prior work commencement.

- PROCEDURES

6.1 Process That Require Validation

General guideline to determine whether a process requires validation.

Routine final product tests have insufficient sensitivity to verify the desired safety and efficacy

Clinical and/or destructive testing would be required to show that the desired results have been achieved. Eg: in the case of sterility where absolute sterility cannot be demonstrated without the complete destruction of every finished article. Refer to Appendix 1 for the Process Validation Decision Tree to determine which processes require validation.

6.2 When to Perform Process Validation / Revalidation

6.2.1 Process validation should be performed whenever there is any new process and/or software set up, new equipment installed for usage in process and new product design resulting in new processes

6.2.2 Re-validation should be performed whenever there is any change stated below but not limited to:

| Type of Changes | Validation to be Conducted |
| --- | --- |
| Transfer of process from one place to another | IQ, OQ, PQ |
| Major modification/change in Hardware or Software | OQ, PQ |
| Change in manufacturing operating condition / parameters | OQ, PQ |
| Negative Trend (s) in quality indicators | OQ, PQ |
| Sterilizer | (OQ, PQ) Annually |
| Change in raw material that could affect machine parameters | OQ, PQ |
| Design change during product realization | OQ, PQ |
| Design change due to post market surveillance | OQ, PQ |

6.3 How to perform Process Validation

6.3.1 Prospective validation

Prospective validation is performed after completion of process design verification. All the requirements on quality, productivity, safety, etc. aspects are expected to be met and process conditions, including process control range are proposed.

- Validations are normally performed in 3 phases (Installation Qualification, Operational Qualification and Performance Qualification).

- Validation includes not only normal operation conditions but also testing upper and lower limits of control condition. In other words, the ‘worst condition’ is considered in the validation to the fullest extent.

- In order to demonstrate process stability and consistency at least 3 runs are required for each process conditions.

- Actual products are used in this type of validation. However, for a non-proven process, similar products or dummy lots can also be used.

- If actual products are used for validation, when validation criteria is achieved and the products comply with all the acceptance requirements, they can be released to the next process or market.

- A validation result is judged according to pre-determined criteria in the approved protocol. After approval of the validation, necessary action for formal implementation of the validated process is taken

The following are the steps for conducting validation:

- Form a validation team, where applicable.

- Plan the approach and define the requirements

- Identify and describe the processes

- Specify process parameters and desired output (acceptance criteria)

- Decide on verification and/or validation

- Select methods and tools for validation

- Create validation protocol

- Perform IQ, OQ, PQ and document reports

- Determine continuous process controls

- Review and update Process FMEA, if necessary

- Control the process continuously

6.3.2 Retrospective Validation

If prospective validation has not been performed in the past and when proof of process adequacy is necessary, retrospective validation can be performed.

To prove the adequacy of a process, inspection and test records and other necessary records shall be collected and analyzed using statistical techniques for its adequacy

6.3.3 Installation Qualification (IQ) Phase

      - Process equipment and ancillary system should be installed, reviewed and calibrated in accordance to manufacturer/supplier approved specification.

      - Process equipment and ancillary system should be evaluated to ensure that it is capable of operating within established limits and tolerances throughout all anticipated operating ranges.

      - Although in most cases, equipment is tested by the suppliers’ prior to delivery, it is usually insufficient to reply entirely on such validation data from suppliers. We must still evaluate and validate that the equipment purchased is suitable for our own usage for the production of out end-product.

      - IQ Documentation

Requirement / documentation for IQ (not limited to):

        - Equipment design

        - Supplier documentation (technical drawings/instruction manuals)

        - Supplier qualification study report (if qualification is done by the supplier prior delivery)

        - Installation requirements (wiring, utilities, etc)

        - Identify important elements of the equipment that could affect output or finished device

        - Safety risk assessment

        - Acceptance criteria

6.3.4 Operational Qualification (OQ) Phase

a) The process parameters are challenged to assure that the end-product meets all the pre-determined specifications and requirements under all anticipated manufacturing conditions – worst case testing.

b) OQ Documentation

Requirement / documentation for OQ (not limited to):

        - Process control limits and operating specification (time, temperature, pressure, speed, etc)

        - Process change control

        - Test items and methods

        - Justification for sample size

        - Acceptance Criteria

c) Justification has to be stated in the validation if worst case conditions are not considered and/or carried out.

6.3.5 Performance Qualification (PQ) Phase

a) To determine that the process consistently produces acceptable end-product under established operating conditions, at least 3 consecutive production lots or 3 consecutive work shifts should be used for PQ to ensure that all external influences can be covered. These influences which can have an impact on the process should also be documented in the protocol.

b) Run Size of each PQ

The run size of each PQ batch should represent the batch size of routine production.

- For process that involve heat distribution (e.g. oven), both minimum and maximum batch size will be validated as amount of load within the process could possibly affect the heat distribution.

- For other processes, maximum batch size should be validated. In the event that there are difficulties to schedule a single lot to meet maximum batch size, multiple batches within a work shift can be used as one PQ run, as long as the process parameters remain unchanged, sampling point are spread and cover every single batch within the PQ run.

      - PQ Documentation

Requirement / documentation for PQ (not limited to):

        - Measurement and monitoring program.

        - Test items and methods

        - Justification for sample size

        - Acceptance Criteria

d) If defect samples were found during the PQ stage, justification has to be provided. This justification has to include that the found defects can also be found during the routine production.

6.4 Validation Protocol and Report

6.4.1 The validation test protocol (FM-07001) enables that the study/program is carefully planned and documented so that at the end of the execution, the process is adequately validated.

The reference number assigned for each validation protocol shall be as followed:

| TP/ | DD/ | XX/ | XXXX |
| --- | --- | --- | --- |
| Test Protocol | Department code | Where XX is a 2 digit running number from 01 to 99 | Where XXXX is a 4 digit running number to indicate the year. |

Elements to be included in the protocol (if necessary):

        - Objective

        - Scope

        - Background

        - Responsibilities

        - Location

        - Reference documents

        - Procedure

        - Acceptance criteria

        - Statistic techniques and sampling sizes

        - Risk assessment

        - Conclusion

        - Re-qualification criteria

        - List of attachment

        - Deviation log and summary

        - Engineering change control and approval

6.4.2 Validation Report

Reference to the protocol number assigned for each validation protocol shall be in accordance with the test protocol as followed;

| TR/ | DD/ | XX/ | XXXX |
| --- | --- | --- | --- |
| Test Report | Department code | Where XX is a 2 digit running number from 01 to 99 | Where XXXX is a 4 digit running number to indicate the year. |

The validation test report (FM-07002) should provide results to all the elements / factors stated in the validation protocol.

  - 6.5   Key and Special Processes

Key and Special processes require increased process monitoring or monitoring of operator performance, as required in order to provide highest confidence in the process performance.
Process monitoring can be accomplished by:

    - Equipment and process parameter monitoring (e.g., setup verification, parameter record) or control (e.g., fixed or limited setting adjustment).

    - Inspection, test and audit methods to verify that the output from the validated process meets specified requirements.  For instance, destructive sampling, using worst-case conditions when possible.

    - For sterilization processes, Company shall follow external standards (if applicable). The validation report shall generate in accordance to the requirements stated in the external standards.

6.6 Sampling Plan

6.6.1 Determine the Criticality

a) Determine the criticality of each product characteristics that will be challenged as acceptance criteria for the validation. The criticality of the product characteristics are determined by the severity of the effect of the failure modes associated with the characteristics being out of specification. The criticality should be defined as one of the three categories: Critical, Major, Minor.

b) Severity of effects from product characteristics failure modes can be found on the relevant “Design Failure Mode Effect Analysis” and “Process Failure Mode Effect Analysis”. Below table maps the 3 criticality categories with the severity table in the “Design Risk Assessment” and “Process Failure Mode Effect Analysis” with reference to QS-06-QA (Risk Management Procedure).

| Criticality of Product Characteristics | Ranking in Design Failure Mode Effect Analysis | Ranking in Process Failure Mode Effect Analysis |
| --- | --- | --- |
| Critical | 4 – 5 | 4 – 5 |
| Major | 2-3 | 2-3 |
| Minor | 1 | 1 |

c) In the event where it is unsure to determine the correct criticality of Product Characteristics, “Critical” should be used.

6.6.2 Determine the Quality Level

For each parameter under study, use the defined Criticality Level (Critical, Major or Minor) associated with each parameter and the type of qualification activity (OQ or PQ), select an appropriate minimum acceptable quality level of Lot Tolerance Percent Defect (LTPD) from below table.

| Qualification Type | Confidence Level (CL) | Criticality Level | LTPD |
| --- | --- | --- | --- |
| OQ | 90% | Critical, Major, Minor | 10% |
| PQ | 90% | Minor | 10% |
|  |  | Major | 5% |
|  |  | Critical | 3% |

6.6.3 Determine the Sampling Size

#### The number of samples is derived using below binomial probability formula with the assumption that the qualification runs simulate continuous stream of production with random sampling:

Below table lists the number of samples based on this equation.

| CL | LTPD | Sample Size (n) and Acceptable Defect Rate (a) |
| --- | --- | --- |
| 90% | 10% | n=22, a=0 |
|  | 5% | n=45, a=0 |
|  | 3% | n=76, a=0 |

6.7 Risk Assessment

Hazards associated with the process being validated should be identified during the course of process validation. Method of risk assessment will be referenced to Risk Management Procedure (QS-06-QA), Process Failure Method Effect Analysis (Process FMEA).

Where process is new or new risk is identified during process validation, the Process FMEA shall be reviewed and updated accordingly.

6.8 Site Validation Master Plan

The Site Validation Master Plan (SVMP) maybe owned by Manufacturing / Engineering Dept.

6.8.1 The SVMP is a document that is to be updated continually to reflect the processes added, relocated/ otherwise changed, or retired.

    - All processes at a location shall be keyed in the Site Validation Master Plan after its pre-implementation validation.

- RECORD

7.1 FM-07001  Process Validation Test Protocol

7.2 FM-07002  Process Validation Test Report

All validation documents are kept as long as the devices is still available on market, plus 2 years.

- APPENDIX

Appendix 1 Process Validation Decision Tree

Appendix 1: Process Validation Decision Tree

The following model shall be used in determining whether or not a process should be validated:

The model shown describes a decision tree that the Company follow when deciding on whether a process needs to be validated.

Each process shall have a specification describing both the process parameters and the output desired. Company shall consider whether the output can be verified by subsequent monitoring or measurement (A). If the answer is positive, then the consideration should be made as to whether or not verification alone is sufficient to eliminate unacceptable risk and is a cost effective solution (B). If yes, the output should be verified and the process should be appropriately controlled (C). If the output of the process is not verifiable then the decision should be to validate the process (D); alternatively, it may become apparent that the product or process should be redesigned to reduce variation and improve the product or process (E). Also, a change in a manufacturing process may result in the need for process validation even though the process formerly only required verification and control.

The risk or cost may also be reduced by redesigning the product or process to a point where simple verification is an acceptable decision (E).