| Roles | Name | Designation | Signature | Date |
| --- | --- | --- | --- | --- |
| Author |  |  |  |  |
| Reviewer |  |  |  |  |
| Approver |  |  |  |  |

| DCO no. | Revision | Change Description | Author | Effective Date |
| --- | --- | --- | --- | --- |
| DCO XXXXX | 01 | Initial Release |  | XX-XXX-XXXX |

- PURPOSE

This procedure is intended to provide a format for software validation and apply to any application that is used in the Quality Management System at Kyron .

- SCOPE

Apply to software used in monitoring, tracking and traceability of material/device used for producing medical device manufactured and distributed by Kyron.

- Self-developed applications

- Server security and user administration

- REFERENCES

  - External Documents

    - ISO 13485:2016

  - Related Documents

    - Quality Manual

4.0  DEFINITIONS

| Commercial Off-the-Shelf (COTS) Software | A generally available software component for which the user (COTS) Software cannot claim complete software life cycle control. It is the readily available software installed and used without any modifications. |
| --- | --- |
| Electronic Records | Any combination of text, graphics, data, audio, pictorial, or other information representation in digital from that is created, modified, maintained, archived, retrieved, or distributed by a computer system. |
| Software | Computer instructions that when executed will provide desired functions and performance, data structures that enable programs to manipulate information and documentation that describes the operation and the use of programs |
| Software Validation | Confirmation by examination and provision of objective evidence that software specifications conform to user needs and intended uses and that the particular requirements implemented through software can be consistently fulfilled. |
| Software Verification | Confirmation by objective evidence that the design outputs of a particular phase of the software development life cycle meet all of the specified input requirements for that phase. |
| Installation Qualification (IQ) | Establishing by objective evidence that all key aspects of the process equipment and ancillary system installation adhere to the manufacturer’s approved specification and that the recommendations of the supplier of the equipment are suitably considered. |
| Operation Qualification (OQ) | Establishing by objective evidence process control limits and action levels which result in product that meets all predetermined requirements |
| Performance Qualification (PQ) | Establishing by objective evidence that the process, under anticipated conditions, consistently produces a product which meets all predetermined requirements. |
| SaMD | Software intended to be used for one or more medical purposes that perform these purposes without being part of a hardware medical device by FDA. |

.

5.0 RESPONSIBILITIES

IT, QA and Manufacturing Engineer are responsible to:

- Draft software validation protocols,

- Conduct software validation and establish report prior to initial use and

- Conduct revalidation, as appropriate, after changes to such software or its application

QA Department Head is to approve the software validation activities.

6.0 PROCEDURE

6.1	Management Review Meeting

Prior to the use of any software, the intended use and requirements of the software should be documented.

A risk analysis should be performed to determine the level of risk should the software not perform as intended.

Top management shall determine during management reviews or other departmental meetings if software to be implemented at Kyron require validation. The conclusion of such meetings shall be retained and filed together with Software List.

6.2	Software Tools Used

Software used in Kyron that contains any of the following functions (non-exhaustive) where such functions are applicable to Kyron ’ medical device or QMS:

•	Developing tools for Software as a Medical Device (SaMD)

•	Database management systems

•	Web applications development platform

•	Document control management systems

•	Microsoft office applications

•	Installation and/or servicing tracking

•	Feedback and complaint handling

•	Personnel training

6.3	Software Tools Used Classification

The results of the risk analysis determine how extensive the software validation should be. If the risk of using the software is high, then an extensive validation approach would be needed. At minimum, the validation should show whether the software is capable of meeting the intended use and requirements.

Kyron software used will have the following classification:

| Software Risk Classification | Definition |
| --- | --- |
| Low | Minimal impact to safety and performance to the medical device e.g. excel spreadsheets used for keeping track of master lists where manual verification is being performed. |
| Medium | Medium impact to safety and performance to the medical device e.g. database management systems/ web application development platforms used by SaMD, software to track training attended by employees or spreadsheets with formula, macros, equations or hyperlinks. |
| High | High impact to safety and performance to the medical device e.g. tracking software used to track expiration dates, customer number creation, delivery or receipts, status of installation and/servicing etc. |

###### 6.4 Validation Requirements

| Software Risk Classification | Validation Requirements |
| --- | --- |
| Low | None required |
| Medium | Minimally verification and acceptability of the functionality of the software. Report shall be maintained |
| High | Perform IQ,OQ & PQ according to Kyron requirements and document the overall acceptability of the software. Report shall be maintained. |

6.4.1 Performing IQ, OQ, PQ

a) Installation Qualification (IQ)

- Purpose

- Utilities

- Test Procedure

- Acceptance Criteria

- Test Result

b) Operation Qualification (OQ)

- Purpose

- Test Procedure

- Acceptance Criteria

- Test Result

c) Performance Qualification (PQ)

- Purpose

- Test Procedure

- Acceptance Criteria

- Test Result

- Revalidation Criteria

6.5 All software from low to high-risk shall be listed in Software List and shall include the version approved for use.

7.0  RECORD

NA

8.0  APPENDIX

NA