| Roles | Name | Designation | Signature | Date |
| --- | --- | --- | --- | --- |
| Author |  |  |  |  |
| Reviewer |  |  |  |  |
| Approver |  |  |  |  |

| Revision History |  |  |  |  |
| --- | --- | --- | --- | --- |
| DCO no. | Revision | Change Description | Author | Effective Date |
| DCO XXXXX | 01 | Initial Release |  | XX-XXX-XXXX |

- PURPOSE

1.1  Define a control system for the identification, documentation, segregation, evaluation and disposition of non-conforming product to prevent its unintended used or delivery.

1.2 To specify the procedure for the department concerned to request and get approval for use of goods on the quality of conformance which did not meet the specified requirements.

- SCOPE

This procedure applies to receiving material, work-in-progress and finished products that are determined to be non-conforming or suspected to be non-conforming.

- REFERENCES

  - External Documents

    - ISO 13485:2016

  - Related Documents

    - Quality Manual

    - QS-06-QA   Risk Management Procedure

    - QS-15-QA   Corrective and Preventive Actions Procedure

    - QS-27-QA   Adverse Event Reporting and Field Safety Corrective Action

4.0   DEFINITIONS

| Non-conforming Product | Incoming, in-process, or finished product that does not conform to Kyron requirements or regulatory requirements. |
| --- | --- |
| Correction | Action to eliminate a detected nonconformity |
| Corrective Action | Action to eliminate the cause of nonconformities in order to prevent recurrence or action undertaken to eliminate the cause of a detected nonconformity or other undesired situation |
| Rework | Action taken on a nonconforming product so that it will fulfil the specified requirements before it is released for distribution |
| Conditional Approved | Conditional to use of parts on the quality of conformance which is not satisfactory and as a special treatment for salvation. |
| NCR | Non-Conformance Report |
| Use As Is | Conditional approved to accept the nonconforming product as it is. |
| Sort | Separate the conforming product from the nonconforming product. |
| QA HOD | Quality Assurance Head Of Department |

5.0    RESPONSIBILITY

5.1 It is the responsibility of all employees to bring to the attention of the QA Department when there is a nonconforming product or suspected nonconforming product found. Any employee can initiate a Non-conformance Report (NCR).

5.2 The QA Department shall control nonconforming product, preventing the inadvertent use   or dispatch of nonconforming product.

5.3 The QA Department is responsible to issue NCR number and update the Non-Conformance Log (FM-12004)

NCR YYYY-XX

where

YYYY: Current Year

XX: Running number

5.4 The QA Department is responsible to issue Rework Instruction number.

RW YYYY-XX

where

YYYY: Current Year

XX: Running number

5.5 Managing Director is responsible to approve conditional approved report.

Conditional Approval (CA) Number and QA to update the Conditional Approved Log (FM-12005)

CA YYYY-XX

where

YYYY: Current Year

XX: Running Number

6.0     PROCEDURES

6.1 Control of Non-Conforming Product (NCP)

6.1.1 All NCP shall be identified with NCR number and segregated at the designated Quarantine Area to prevent inadvertent use or delivery to customer.

6.1.2 NCP shall be traced and the possible affected NCP scope shall be identified, segregated and put on hold. Containment action for current production shall be taken. The entire batch or shipment may be quarantined pending investigation to confirm the status of the product.

6.1.3 Root cause and corrective action is required for each NCR raised.

6.1.4 A discussion/risk assessment shall be held with the concerns department to confirm the disposition:

- Rework, Use as Is, Scrap, Return to Supplier & Sorting for products that is in Kyron inventory.

- Recall for products or Use as Is for products that have been shipped out.

- The detail of the disposition is specified and approved by the QA HOD.

6.1.5 For finished product to be scrapped, the Write-Off Requisition form (FM-33003) shall be issued and submitted to respective department heads for confirmation, verification and approval before disposition is carried out.

6.1.6 NCP detected after delivery or use shall be assessed with the concerned departments to confirm the next action to be taken, depending on the nature and condition of the defect. Action taken shall be commensurate with risk, or potential risks, of the non-conformity.

6.1.7 For NCP that have been shipped out and subjected to recall, refer to Adverse Event Reporting and Field Safety Corrective Procedure (QS-27-RA).

6.1.8 Identification of the root cause and corrective actions to be taken in accordance to Corrective and Preventive Actions Procedure (QS-15-QA).

6.1.9 Risk assessment shall be completed and approved by QA HOD.

6.2 Conditional Approval Control

6.2.1 Department requesting for conditional approval shall initiate the Conditional Approval Report (FM-12003).

6.2.2 UAI could only be granted when there is no quality, safety and performance impact that could affect the compliance of the devices against the applicable regulatory requirements

6.2.2 All conditional approval requests shall be final approved by Managing Director.

6.3 Product Rework Control

6.3.1 Rework must be performed in accordance with documented Rework Instruction (FM-12002) taking into account the potential adverse effect of the rework on the product.

6.3.2 Ensure there’s no safety and performance impact that could affect the compliance of the devices against the applicable regulatory requirements.

6.3.3 Assign a new lot/batch number for the reworked products.

6.3.4 The rework instruction must be approved by the same functions which approved the original work instructions.

6.3.5 Upon approval on the rework instruction, perform rework according to the rework instructions. Identification of the operator performing rework shall be recorded.

6.3.6 All reworked products shall be re-verified or re-inspected by QA department and recorded in NCR.

6.3.7 Records of rework shall be maintained as part of the Device History Records (DHR).

6.4 NCR Closure

QA Department shall verify that the disposition of NCP has been carried out accordingly and close the NCR.

6.5 Corrective Action Request

Requirement to raise a CAR for each non-conformance will be assess by the QA HOD based on the severity and occurrence rate of the reported non-conformance.

7.0    RECORD

7.1 FM-12001   Non-Conformance Report

7.2 FM-12002   Rework Instruction

7.3 FM-12003   Conditional Approved Report

7.4 FM-12004   Non-Conformance Log

7.5 FM-12005   Conditional Approved Log

8.0    APPENDIX

NA