| Roles | Name | Designation | Signature | Date |
| --- | --- | --- | --- | --- |
| Author |  |  |  |  |
| Reviewer |  |  |  |  |
| Approver |  |  |  |  |

| Revision History |  |  |  |  |
| --- | --- | --- | --- | --- |
| DCO no. | Revision | Change Description | Author | Effective Date |
| DCO XXXXX | 01 | Initial Release |  | XX-XXX-XXXX |

- PURPOSE

  - To define the system for identification and traceability requirements for purchased and manufactured parts.

  - To ensure that status for raw materials, in-process materials and finished products are identified during all stages of receipt, production, storage, distribution, installation, and servicing.

- SCOPE

This procedure applies to all raw materials, sub-assemblies, and finished goods purchased, or manufactured by Kyron.

- REFERENCES

  - External Documents

    - ISO 13485:2016

  - Related Documents

3.2.1  Quality Manual

3.2.2  QS-12-QA   Control of Non-conforming Product Procedure

4.0    DEFINITIONS

| Lot | One or more components or finished devices that consist of a single type, model, class, size, composition, or software version that are manufactured under essentially the same conditions and that are intended to have uniform characteristics and quality within specified limits. |
| --- | --- |
| Lot / Batch Number | A unique number identifying a particular batch material, sub-assemblies, or finished goods. |
| Serial Number | A unique number identifying a single unit of raw material, sub-assembly, or finished good. |
| Traceability | Ability to trace the history, application or location of that which is under consideration. |
| Identification | A method of allowing identification status of products or unique code on the product. |
| Device History Record (DHR) | The relevant records of each batch/lot or unit of the medical device that is manufactured from raw materials to finished product, including identification of measuring equipment. |

5.0     RESPONSIBILITY

5.1	Quality Assurance (QA) Department is responsible to

- Assign incoming in-house batch number for raw material.

- For archiving all DHR records and distribution records.

5.2	QA Department shall liaise with Production Manager/Engineer on product coding identification implementation for sub-assemblies and finished goods.

5.3	QA Department shall work together with Production/Supply Chain/Sales & Marketing staffs for product traceability.

6.0     PROCEDURES

6.1	Identification and Traceability of Purchased Raw Material

6.1.1 All incoming purchased material shall be staged at “Quarantine Area” pending for inspection prior to receipt into warehouse inventory.

6.1.2 Purchased material are identified with unique part number and lot number. In-house lot number may be assigned.

6.1.3 The lot number is assigned according to the format below:

- YYMMDDAABB where:

YYMMDD represent the date (Year, Month, Day) (e.g. 200408)
AA represent material

BB represent running number

Material Code:

| AA | AB | AC | AD |
| --- | --- | --- | --- |
| Plastic tube |  |  | Packaging box |

6.1.4 Acceptance Status:

6.1.4.1 Material that passed incoming inspection will be labelled as “Approved” and shall move to designated location in warehouse approved for use.

6.1.4.2 Material that failed incoming inspection will be labelled as “Quarantine” and shall follow the Control of Non-conforming Product procedure (QS-12-QA) to manage these non-conforming products.

6.1.5 All materials are labelled and identified with the unique part numbers, lot numbers and acceptance status to prevent mix-ups and to allow effective traceability of the product. Material may be identified using barcodes.  Barcodes contain the information necessary to identify product, part and lot number information.

6.1.6 Production shall use only materials that are approved for use.

6.2	Identification and Traceability During Production

6.2.1 During all stages of production, manufactured parts and sub-assemblies are identified by the part number and lot number on the batch record. They may be identified with suitable indicator on the bin or containers holding the parts or the numbers may be marked directly on the parts.

6.2.2 Parts/materials lot/batch number that was used on the device shall be recorded as part of traceability requirements

6.2.3 Acceptance Status:

6.2.2.1 Manufactured parts or sub-assemblies that passed in-process inspection will be released for the next steps of the manufacturing process. The product status is indicated on the batch record.

6.2.2.2 Manufactured parts or sub-assemblies that failed in-process inspection will be labelled as “Quarantine” or place at “Quarantine” area and may follow-up with the Control of Non-conforming Product Procedure to manage these non-conforming products.

6.3	Identification and Traceability of Finished Product

6.3.1 Acceptance Status:

6.3.1.1 Finished products that passed final inspection will be labelled with product label and authorized for release to customers.

6.3.1.2 Finished products that failed final inspection will be labelled as “Quarantine” or place at “Quarantine” area and follow-up with the Control of Non-conforming Product Procedure to manage these non-conforming products.

6.3.2 Finished products are identified by product label that is affixed to the product or its packaging.

6.3.3 The lot number is assigned to the batch of production. The lot number is assigned according to the format below:

  - PYYYYMMXXX where:

P: represent Finished Product Code

YYYYMM001: represent the date (Year, Month)

XXX represent Batch code

(E.g. AHC202210001)

Finished Product Code

| Product name | Model number | Product code | Lot number (Eg) |
| --- | --- | --- | --- |
| Annuflow-HC | AHC-1000 | AHC | AHC202210001 |

6.3.4 The finished products are kept at the designated proper storage location prior to distribution.

6.4	Expiry Date

Expiry date of product shall be written as:

DD-MM-YYYY

(Eg. 18-08- 2025)

6.5	Contract Manufacturing

For contract manufacturing, the device history record shall refer to the manufacturer’s QMS for traceability.

6.6	Device History Record and Traceability

The Device History Record shall contain all the information related in manufacturing of the medical device to allow product traceability. It shall contain but not limited to;

- Critical raw material identification

- Work order number

- Product information

- QA and production inspection records

- Equipment identification

- Packaging/labelling information (lot/batch number, expiry, etc)

- Approved device quantity to be shipped

.

7.0	RECORD

7.1 FM-16001   Incoming Material Lot Number Assignment Log

8.0	APPENDIX

NA